COMPANION TOOLKIT · TOOL T-1

# Device Signal-Management SOP Outline and RACI

Structure for a controlled signal-management procedure, with a RACI aligned with ISO 13485 processes

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Gives the section structure for a controlled device signal-management procedure and a RACI that assigns one accountable owner to each step, so that a question raised in complaints, CAPA or risk management cannot leak between processes. |
| --- | --- |
| When to use | When writing or revising the signal-management SOP; when a CAPA, complaint or risk-file review shows that nobody owned the whole question; before a quality-system audit. |
| Chapter reference | Chapter 4, Building the Signal System Inside the Quality System — section 4.5 (Tables 4.3 and 4.4). |

# Document control

| SOP number |  | Version |  |
| --- | --- | --- | --- |
| Owner (function) |  | Effective date |  |
| Approved by |  | Next review due |  |
| Devices / markets in scope |  | Related SOPs (complaints, CAPA, risk, vigilance reporting) |  |

# Part A — SOP outline

Use the twelve sections below. The last column records where each element sits in your own procedure. Both parts are examples to adapt, not regulatory templates.

| SOP section | Content | Key record | Your SOP section / record ID | Owner |
| --- | --- | --- | --- | --- |
| 1. Purpose and scope | Devices, markets and sources covered | — |  |  |
| 2. Definitions | Controlled terms from Chapter 2 | Glossary |  |  |
| 3. Roles and governance | Owners, review board, links to CAPA board and risk team | RACI |  |  |
| 4. Detection and intake | Sources, screening methods, route for non-complaint information | Data-source inventory (Tool T-2) |  |  |
| 5. Opening a signal record | Criteria, required fields, scope | Signal record |  |  |
| 6. Validation and prioritization | Criteria and timelines set by the company | Validation memo |  |  |
| 7. Assessment | Protocol, clinical and engineering review | Assessment report |  |  |
| 8. Decision and action | Review board authority, escalation, links to CAPA and FSCA | Decision minutes |  |  |
| 9. Risk file and document updates | Inputs to ISO 14971 clause 10 review; PMS deliverables | Risk file update record |  |  |
| 10. Closure and reopening | Conclusion categories, rationale, reopening criteria | Closure memo |  |  |
| 11. Urgent escalation | Same-day route for serious clustered events | Escalation log |  |  |
| 12. Metrics and review | Paired speed and quality metrics | KPI report |  |  |

Outline as shown in Table 4.3 of the book (example).

## Minimum requirements to write into the SOP

☐  Before any CAPA closes, a search for related records by device, component and failure mode is performed and its result recorded.

☐  One function is accountable for opening, validating and concluding signals, with authority to call on the others.

☐  CAPA closures and complaint trends are named inputs to the ISO 14971 clause 10 production and post-production review.

☐  Conclusions use the controlled categories confirmed, refuted or indeterminate; "closed" is an administrative status, not a finding.

☐  Every signal record states its reopening criteria.

☐  Company-designed thresholds and timelines are labeled as company choices, never as regulatory requirements.

# Part B — RACI aligned with ISO 13485 processes

R = responsible (does the work); A = accountable (one function only, signs off); C = consulted; I = informed. The book's example shows R, A and C; add the informed parties for your organization.

| Process step | ISO 13485 anchor | Responsible | Accountable | Consulted | Informed |
| --- | --- | --- | --- | --- | --- |
| Complaint intake and linkage search | 8.2.2 |  |  |  |  |
| Feedback and non-complaint sources | 8.2.1 |  |  |  |  |
| Signal record opening and validation | 8.2.1 |  |  |  |  |
| Reporting to authorities | 8.2.3 |  |  |  |  |
| Corrective action | 8.5.2 |  |  |  |  |
| Preventive action | 8.5.3 |  |  |  |  |
| Signal conclusion and action | 8.2.1, 8.5.2 |  |  |  |  |

If two functions claim accountability for the same step, resolve it before the SOP is approved.

### Approval

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| SOP owner |  |  |  |
| Head of quality |  |  |  |
| Vigilance lead |  |  |  |
| Regulatory lead |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Table 4.4 of the book: RACI partly filled; the last row is left for the reader's organization.

| Process step | ISO 13485 anchor | Responsible | Accountable | Consulted |
| --- | --- | --- | --- | --- |
| Complaint intake and linkage search | 8.2.2 | Complaint handling | Complaint-handling lead | Vigilance |
| Feedback and non-complaint sources | 8.2.1 | PMS | PMS lead | Service; supplier quality; sales |
| Signal record opening and validation | 8.2.1 | Vigilance | Vigilance lead | Clinical; R&D |
| Reporting to authorities | 8.2.3 | Regulatory affairs | Regulatory lead | Vigilance |
| Corrective action | 8.5.2 | CAPA owner | Head of quality | Vigilance; R&D |
| Preventive action | 8.5.3 | CAPA owner | Head of quality | Risk management |
| Signal conclusion and action | 8.2.1, 8.5.2 | Review board | [to be assigned] | [to be assigned] |

Table 4.4 — Tool T-1: RACI aligned with ISO 13485 processes (partly filled example). Kestran context: three insulation-abrasion CAPAs at three hospitals were closed without being linked, because no step owned the linkage search.