## Read me
Time-to-failure analysis plan and Kaplan–Meier template
Signal Management in Medical Device Vigilance · Selim Medical Press · © 2026 · Tool T-8 · Chapter 10 (section 10.5; Tables 10.1–10.2)
Educational template — adapt to your quality system. Not a regulatory requirement. All example data are fictional.
Purpose
The Plan sheet fixes the question, time zero, censoring, strata, method and review trigger before the analysis runs. The life-table sheets take intervals, devices at risk, failures and censored counts, and return interval hazard, Kaplan–Meier survival and Greenwood linear 95% confidence intervals, with charts of the curve and the hazard. The strata sheets repeat the calculation for each failure mode, model, lot or period.
When to use
For implants and other long-lived devices where failures accrue over time and a crude proportion (failures ÷ devices) understates cumulative risk because younger devices have not reached the high-hazard years.
Sheets
Plan — Table 10.2 filled for Kestran, with a column for your entries and an approval block.
KM Example — Kestran (fictional): fractures 3, 4, 5, 14, 22; at risk 4,000, 4,000, 3,200, 2,400, 1,600. Reproduces 5-year survival 97.7% (95% CI 97.1–98.4).
KM Your data — blank, 10 intervals. Strata Example — Kestran header vs conductor fractures. Strata Your data — three blank strata.
Formulas
Hazard = d ÷ n. S(t) = S(t−1) × (1 − d ÷ n). Greenwood sum = Σ d ÷ [n × (n − d)]. SE = S × √(Greenwood sum). 95% CI = S ± z × SE, truncated to 0–100%.
Caution
Review triggers are company-designed examples. Where death is common, Kaplan–Meier overstates long-term failure; treat death as a competing risk as the plan specifies.
Companion web tool: https://abdu94-hash.github.io/device-signal-tool/

## Plan
Time-to-failure analysis plan — complete and approve BEFORE the analysis runs
Signal Management in Medical Device Vigilance · Selim Medical Press · © 2026 · Tool T-8 · Chapter 10 (Table 10.2)
Educational template — adapt to your quality system. Not a regulatory requirement. All example data are fictional.
Device / model:
Signal record ID:
Data cut:
Plan element | Kestran entry (example, fictional) | Your entry
Question and decision it informs | 5-year fracture risk by mode; informs physician advisory, follow-up interval and design action
Time zero and imputation rule | Implant date; missing dates imputed from ship date plus documented lag, with sensitivity analysis
Censoring and competing risks | Censor at data cut, loss to follow-up and non-fracture explant; death as competing risk; suspected-fracture explants adjudicated first
Strata | Failure location (header, conductor at subclavian crossing); implant year; predecessor lead
Method and outputs | Kaplan–Meier at annual intervals, Greenwood linear 95% CI, Weibull shape by mode; number at risk shown at each year
Pre-set review trigger (company example) | Annual hazard in any year above the predecessor's, or any new failure location in returned leads
The review trigger is a company-designed example. Its justification belongs in the PMS plan (Tool T-5), with the observation period stated.
Approval (before any data are analyzed)
Role | Name | Signature and date
Signal owner
Statistician
Clinical lead
Amendments after approval (report any post hoc stratum as exploratory)
Date | Change and reason | Approved by

## KM Example
Kaplan–Meier life table with Greenwood 95% CI — Example (Kestran, fictional)
Signal Management in Medical Device Vigilance · Selim Medical Press · © 2026 · Tool T-8 · Chapter 10
Educational template — adapt to your quality system. Not a regulatory requirement. All example data are fictional.
Legend: | Input (edit) | Calculated (do not edit)
Confidence level |  | 95% | [12 empty] | Enter at-risk counts at the start of each interval directly (after censoring and earlier failures). Censored counts are recorded for traceability only. The linear CI is truncated at 0% and 100%.
z | [14 empty] | Grey chart-helper columns (O–R) deliberately return the N/A error value for empty intervals so the charts leave gaps.
Cohort and data cut |  | Tracked cohort of 4,000 leads (800 implanted in each of Years 1–5); data cut end of Year 6
LIFE TABLE
Interval (e.g., year since implant) | At risk at start (n) | Failures (d) | Censored (c) — record only | Interval hazard d ÷ n | Conditional survival 1 − d/n | Kaplan–Meier survival S | Cumulative failure 1 − S | Greenwood term d ÷ [n(n − d)] | Greenwood sum | SE of S | Lower CI (linear) | Upper CI (linear) | Truncated at 0 or 100%? | chart: S | chart: lower | chart: upper | chart: hazard
1 | 4000 | 3
2 | 4000 | 4
3 | 3200 | 5
4 | 2400 | 14
5 | 1600 | 22
Summary
Crude proportion = total failures ÷ at risk in first interval
Total failures
Survival at last completed interval
lower CI
upper CI
Cumulative failure probability at last interval
Hazard: last interval ÷ first interval
Book reference (Table 10.1): survival 99.93%, 99.83%, 99.67%, 99.09%, 97.7% (97.1–98.4); hazard 0.08% → 1.38%; crude 1.2%; cumulative fracture probability 2.3%. Predecessor lead 5-year survival 99.2% lies above the upper limit.
Censored counts are not given in the case dossier; the book's at-risk counts are entered directly.
Result paragraph for the signal record: "In the tracked cohort of 4,000 leads (data cut end of Year 6), Kaplan–Meier fracture-free survival at 5 years is 97.7% (95% CI 97.1–98.4; 1,600 leads at risk in year 5), against 99.2% for the predecessor lead. Annual hazard rises from 0.08% in year 1 to 1.38% in year 5. Two modes are present: conductor fracture at the subclavian crossing (41 events, wear-out) and header fracture (7 events, all in years 1 to 3)."

## KM Your data
Kaplan–Meier life table with Greenwood 95% CI — Your data
Signal Management in Medical Device Vigilance · Selim Medical Press · © 2026 · Tool T-8 · Chapter 10
Educational template — adapt to your quality system. Not a regulatory requirement. All example data are fictional.
Legend: | Input (edit) | Calculated (do not edit)
Confidence level |  | 95% | [12 empty] | Enter at-risk counts at the start of each interval directly (after censoring and earlier failures). Censored counts are recorded for traceability only. The linear CI is truncated at 0% and 100%.
z | [14 empty] | Grey chart-helper columns (O–R) deliberately return the N/A error value for empty intervals so the charts leave gaps.
Cohort and data cut
LIFE TABLE
Interval (e.g., year since implant) | At risk at start (n) | Failures (d) | Censored (c) — record only | Interval hazard d ÷ n | Conditional survival 1 − d/n | Kaplan–Meier survival S | Cumulative failure 1 − S | Greenwood term d ÷ [n(n − d)] | Greenwood sum | SE of S | Lower CI (linear) | Upper CI (linear) | Truncated at 0 or 100%? | chart: S | chart: lower | chart: upper | chart: hazard
Summary
Crude proportion = total failures ÷ at risk in first interval
Total failures
Survival at last completed interval
lower CI
upper CI
Cumulative failure probability at last interval
Hazard: last interval ÷ first interval

## Strata Example
Stratified life tables (failure mode, model, lot or period) — Example (Kestran, fictional)
Signal Management in Medical Device Vigilance · Selim Medical Press · © 2026 · Tool T-8 · Chapter 10
Educational template — adapt to your quality system. Not a regulatory requirement. All example data are fictional.
Legend: | Input (edit) | Calculated (do not edit)
Confidence level |  | 95% |  | For a failure-mode stratum, count only that mode's failures and treat the other mode as censored. In the example, at-risk counts are those of the pooled table, as in the book (Chapter 10, Step 6).
z
STRATUM 1 — Header fracture (early mode)
Interval (e.g., year since implant) | At risk at start (n) | Failures (d) | Censored (c) — record only | Interval hazard d ÷ n | Conditional survival 1 − d/n | Kaplan–Meier survival S | Cumulative failure 1 − S | Greenwood term d ÷ [n(n − d)] | Greenwood sum | SE of S | Lower CI (linear) | Upper CI (linear) | Truncated at 0 or 100%?
1 | 4000 | 2
2 | 4000 | 2
3 | 3200 | 3
4 | 2400 | 0
5 | 1600 | 0
Summary
Crude proportion = total failures ÷ at risk in first interval
Total failures
Survival at last completed interval
lower CI
upper CI
Cumulative failure probability at last interval
Hazard: last interval ÷ first interval
STRATUM 2 — Conductor fracture at the subclavian crossing (wear-out mode)
Interval (e.g., year since implant) | At risk at start (n) | Failures (d) | Censored (c) — record only | Interval hazard d ÷ n | Conditional survival 1 − d/n | Kaplan–Meier survival S | Cumulative failure 1 − S | Greenwood term d ÷ [n(n − d)] | Greenwood sum | SE of S | Lower CI (linear) | Upper CI (linear) | Truncated at 0 or 100%?
1 | 4000 | 1
2 | 4000 | 2
3 | 3200 | 2
4 | 2400 | 14
5 | 1600 | 22
Summary
Crude proportion = total failures ÷ at risk in first interval
Total failures
Survival at last completed interval
lower CI
upper CI
Cumulative failure probability at last interval
Hazard: last interval ÷ first interval
Book reference: 5-year cumulative probability 0.19% for header fracture and 2.09% for conductor fracture. Weibull shape (fitted outside this workbook): pooled 3.0 (2.3–3.9); conductor 4.3 (3.2–5.6); header 0.8 (0.3–1.7).

## Strata Your data
Stratified life tables (failure mode, model, lot or period) — Your data
Signal Management in Medical Device Vigilance · Selim Medical Press · © 2026 · Tool T-8 · Chapter 10
Educational template — adapt to your quality system. Not a regulatory requirement. All example data are fictional.
Legend: | Input (edit) | Calculated (do not edit)
Confidence level |  | 95% |  | For a failure-mode stratum, count only that mode's failures and treat the other mode as censored. In the example, at-risk counts are those of the pooled table, as in the book (Chapter 10, Step 6).
z
STRATUM 1 — name: ________
Interval (e.g., year since implant) | At risk at start (n) | Failures (d) | Censored (c) — record only | Interval hazard d ÷ n | Conditional survival 1 − d/n | Kaplan–Meier survival S | Cumulative failure 1 − S | Greenwood term d ÷ [n(n − d)] | Greenwood sum | SE of S | Lower CI (linear) | Upper CI (linear) | Truncated at 0 or 100%?
Summary
Crude proportion = total failures ÷ at risk in first interval
Total failures
Survival at last completed interval
lower CI
upper CI
Cumulative failure probability at last interval
Hazard: last interval ÷ first interval
STRATUM 2 — name: ________
Interval (e.g., year since implant) | At risk at start (n) | Failures (d) | Censored (c) — record only | Interval hazard d ÷ n | Conditional survival 1 − d/n | Kaplan–Meier survival S | Cumulative failure 1 − S | Greenwood term d ÷ [n(n − d)] | Greenwood sum | SE of S | Lower CI (linear) | Upper CI (linear) | Truncated at 0 or 100%?
Summary
Crude proportion = total failures ÷ at risk in first interval
Total failures
Survival at last completed interval
lower CI
upper CI
Cumulative failure probability at last interval
Hazard: last interval ÷ first interval
STRATUM 3 — name: ________
Interval (e.g., year since implant) | At risk at start (n) | Failures (d) | Censored (c) — record only | Interval hazard d ÷ n | Conditional survival 1 − d/n | Kaplan–Meier survival S | Cumulative failure 1 − S | Greenwood term d ÷ [n(n − d)] | Greenwood sum | SE of S | Lower CI (linear) | Upper CI (linear) | Truncated at 0 or 100%?
Summary
Crude proportion = total failures ÷ at risk in first interval
Total failures
Survival at last completed interval
lower CI
upper CI
Cumulative failure probability at last interval
Hazard: last interval ÷ first interval