COMPANION TOOLKIT · TOOL T-9

# Special-Situation Checklist and AI Performance-Monitoring Plan

Part A: which special situations apply and whether the system captures them · Part B: monitoring plan for AI devices

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Part A asks, for each product, which special detection situations apply and whether the complaint form, coding rules and data sources capture the extra elements each one needs. Part B defines, for AI devices, what is sampled, how often, how ground truth is established, which strata are reported and which thresholds trigger review. |
| --- | --- |
| When to use | At design transfer and at each PMS plan update; whenever software, firmware, a connected function, a medicinal substance, reprocessing or a third-party accessory is introduced. |
| Chapter reference | Chapter 11, Special Detection Situations — section 11.5 (Tables 11.1 and 11.2). |

# Part A — Special-situation checklist

| Product / model |  | Software version(s) |  |
| --- | --- | --- | --- |
| Completed by |  | Date |  |

| Situation | Applies? (Y/N) | Extra elements captured on complaint form? (Y/N/partly) | Coding rules and data sources in place? | Gap and action (owner, due) |
| --- | --- | --- | --- | --- |
| Software and AI |  |  |  |  |
| Connected and cybersecurity |  |  |  |  |
| Combination product |  |  |  |  |
| Use error |  |  |  |  |
| Reprocessing and accessories |  |  |  |  |

### Reference: what each situation needs

| Situation | Extra data elements | Detection method | Regulatory touchpoint |
| --- | --- | --- | --- |
| Software and AI | Version, configuration, input source, site | Ground-truth sampling, input-drift monitoring, version-level rates | IEC 62304; FDA PCCP guidance |
| Connected and cybersecurity | Firmware, network context, vulnerability ID | Disclosure intake, advisories, telemetry | IEC 81001-5-1 |
| Combination product | Constituent, approach, dose or load, MedDRA term | Constituent-level review, PMCF with standard definitions | MDR Art 1(8), Annex IX section 5.2 |
| Use error | User type, setting, task step, training | Complaint narrative review, simulated use | IEC 62366-1 |
| Reprocessing and accessories | Reprocessing history, accessory source | Site clusters, returned-device evaluation | ISO 13485 clause 8.2.2 |

Table 11.1 of the book — special situations: extra data, detection method and regulatory touchpoint.

# Part B — AI performance-monitoring plan

| AI device / version |  | Intended use |  |
| --- | --- | --- | --- |
| Deployed sites (number) |  | Plan owner |  |

Thresholds are company-designed examples; justify them in the PMS plan.

| Performance metrics and reference |  |
| --- | --- |
| Ground truth and sampling (what, how often, who confirms) |  |
| Strata reported |  |
| Input monitoring |  |
| Review trigger (company example) |  |
| Actions when the trigger fires (change control, user notice, sampling frequency) |  |

### Stratum results log

| Period | Stratum | Estimate (n/N) | 95% CI | Trigger met? (Y/N) |
| --- | --- | --- | --- | --- |

### Approval

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| PMS plan owner |  |  |  |
| Software / AI lead |  |  |  |
| Vigilance lead |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Lumora AI chest X-ray triage software, Part B, Year 2, quarter 3. Zero complaints were received; the drift was found only by sampling.

| Plan element | Lumora entry |
| --- | --- |
| Performance metrics and reference | Sensitivity and specificity against premarket 94% (470/500) and 91% |
| Ground truth and sampling | Radiologist-confirmed positives, quarterly; monthly for two quarters after a confirmed drift |
| Strata | Site; scanner model and firmware; acquisition protocol; software version |
| Input monitoring | Automated image-parameter drift checks at every site; site log of scanner and protocol changes |
| Review trigger (company example) | Any stratum whose sensitivity upper 95% limit falls below the premarket lower limit (91.6%) |

Table 11.2 — Tool T-9 AI performance-monitoring plan, filled for Lumora (illustrative).

| Period | Stratum | Estimate (n/N) | 95% CI | Trigger met? |
| --- | --- | --- | --- | --- |
| Year 2, Q3 | 3 hospitals after scanner firmware update | 86% (129/150) | 79.5–90.7 | Yes (90.7% < 91.6%) |
| Year 2, Q3 | All other sampled sites | 93% (279/300) | 89.5–95.4 | No |

Under this trigger, the affected-site stratum would have prompted review in the quarter it appeared.