COMPANION TOOLKIT · TOOL T-10

# Validation Memo

One-to-two-page record for every alert that reaches validation, whatever the outcome

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Documents how the validation decision was made, by whom and on what evidence. A validated memo becomes the first entry in the signal record; a non-validated memo stays in the signal log with its monitoring plan and reopening trigger. |
| --- | --- |
| When to use | For every alert that reaches validation. Signed by the clinical and engineering reviewers. |
| Chapter reference | Chapter 12, Signal Validation — section 12.5 (Tables 12.1 and 12.2). |

# Identification

| Memo / alert ID |  | Product and model |  |
| --- | --- | --- | --- |
| Data cut |  | Source of the alert |  |

# Part 1 — Criteria

| Criterion | Evidence for | Evidence against or missing | Judgment |
| --- | --- | --- | --- |
| Novelty |  |  |  |
| Relevance |  |  |  |
| Consistency |  |  |  |
| Alternatives |  |  |  |
| Related history |  |  |  |

Judgment examples: Met / Not met / Supports / Open, for assessment / Link in record.

# Part 2 — Memo

| Alert and data cut |  |
| --- | --- |
| Evidence for |  |
| Evidence against and missing |  |
| Decision and scope (validated / not validated; scope of the assessment question) |  |
| Next steps, reviewers, date |  |

A memo that cannot fill "evidence against and missing" has not looked hard enough.

## If not validated

| Monitoring plan |  |
| --- | --- |
| Reopening trigger |  |
| Retained in signal log (reference) |  |

### Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Clinical reviewer |  |  |  |
| Engineering reviewer |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Kestran defibrillator lead, Year 5, quarter 2.

| Criterion | Evidence for | Evidence against or missing | Judgment |
| --- | --- | --- | --- |
| Novelty | New header location; all three under 3 years from implant | Fracture as such is labeled | Met |
| Relevance | Three confirmed fractures on returned-device evaluation | Harm in these three patients not yet documented | Met |
| Consistency | 3 of 3 returned leads at the same location | 6 of 9 leads not returned; location unknown | Supports |
| Alternatives | None yet identified | Center, lot, implant technique, explant damage unchecked | Open, for assessment |
| Related history | Three unlinked insulation-abrasion CAPAs in Years 3–4 | Different failure mode; relationship unknown | Link in record |

Table 12.1 — Kestran validation criteria, Year 5, quarter 2 (illustrative).

| Field | Kestran entry |
| --- | --- |
| Alert and data cut | 9 fracture complaints, Year 5, quarter 2; 3 returned leads with fracture just distal to the header, all implanted under 3 years earlier |
| Evidence for | New failure location; early timing; 3 of 3 returned leads consistent; potentially serious harm |
| Evidence against and missing | Fracture labeled; 6 of 9 leads not returned; center, lot and implant technique not yet reviewed |
| Decision and scope | Validated; assess header fracture as a potential early failure mode; link three earlier insulation-abrasion CAPAs |
| Next steps, reviewers, date | Request center, lot and implant data; recover unreturned leads or imaging; search past returned-device reports; safety physician and design engineer, with date |

Table 12.2 — Tool T-10 validation memo, filled for Kestran (illustrative).