COMPANION TOOLKIT · TOOL T-11

# Prioritization Matrix Aligned with ISO 14971

Anchor table, band rules, floor rules, urgency triggers and a one-page scoring worksheet

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Decides the order and speed of assessment for validated signals using five domains scored 1–5 on the same severity scale as the risk management file. It does not conclude anything about causality and never replaces or extends a legal reporting clock. |
| --- | --- |
| When to use | At validation, and again whenever the data change (a score is a snapshot). Calibrate once a year. Everything here is a company design, offered as an example — no regulation prescribes a device prioritization method. |
| Chapter reference | Chapter 13, Signal Prioritization — section 13.5 (Tables 13.1–13.3). Companion web tool: https://abdu94-hash.github.io/device-signal-tool/ |

# Reference 1 — Scoring anchors (company example)

| Domain | 1 | 3 | 5 |
| --- | --- | --- | --- |
| Severity (risk file scale) | Negligible: no injury or temporary discomfort | Serious: injury needing medical intervention | Catastrophic: death or life-threatening injury credible |
| Probability change | Within the risk file estimate | Estimate exceeded, up to fivefold | Fivefold or more and rising over consecutive periods |
| Exposure | Fewer than 1,000 units in use; not shipping | 1,000 to 10,000 units in use | More than 100,000 units, or an implanted base that cannot be recalled |
| Detectability | Failure obvious before use | Detected by an alarm or routine check, not always in time | Silent until harm; no alarm or cue |
| Vulnerability | Trained professional users, low-risk population | Professional users, high-acuity setting | Population cannot self-rescue: neonates, young children, sedated patients |

Table 13.1 — Example scoring anchors. Scores 2 and 4 sit between the anchors shown; score 4 for probability change is "fivefold or more, trajectory stable or unknown", and score 4 for exposure is "10,001 to 100,000 units in use, affected configuration still distributed".

Severity: score the worst credible harm, not the worst harm reported so far; use the risk file's five levels and definitions.

Probability change: how far the post-market estimate has moved from the risk file estimate, and whether it is still moving (P1 and P2 per ISO/TR 24971).

Exposure: devices in use, implanted base or procedures per year for the affected model, version or lot; is the affected configuration still shipping?

Detectability — direction: a HIGH score means the failure is HARD to detect (the opposite of some FMEA scales). Print this on every form.

Vulnerability: life-sustaining therapy, home use by lay users, populations that cannot self-rescue.

# Reference 2 — Bands, floor rules, urgency triggers, overrides

| Band | Total score | Company response time (example) |
| --- | --- | --- |
| A | 20 to 25 | Board convenes within 5 working days; action decision recorded within 15 calendar days |
| B | 14 to 19 | Assessment plan within 30 days; decision within 90 days |
| C | 5 to 13 | Review at the next scheduled trend meeting |

Company timelines sit beside the legal reporting clocks and never extend them.

Floor rule: any signal with severity 5 and detectability 4 or 5 is placed in Band B at least.

Urgency triggers (any one → Band A regardless of score): a death possibly linked to the device; an enquiry from a competent authority or FDA; a competitor FSCA for the same mechanism; media attention; bench reproduction of a hazardous failure with no available mitigation.

Overrides: the board may move a signal one band up or down. Up needs a written reason. Down needs a written reason and the chair's signature.

# Scoring worksheet

## Part 1 — Signal identification

| Signal record number |  | Device |  |
| --- | --- | --- | --- |
| Model or version |  | Data cut |  |
| Scoring occasion (validation / re-score) |  | Date |  |

## Part 2 — Scores (two independent reviewers)

| Domain | Reviewer 1 | Reviewer 2 | Agreed | Rationale (mandatory) and risk file reference |
| --- | --- | --- | --- | --- |
| Severity |  |  |  |  |
| Probability change |  |  |  |  |
| Exposure |  |  |  |  |
| Detectability (high = hard to detect) |  |  |  |  |
| Vulnerability |  |  |  |  |
| Total (5–25) |  |  |  |  |

Any domain where the reviewers differ by more than one point is discussed and the agreed value recorded.

## Part 3 — Rules

| Total and band |  |
| --- | --- |
| Floor rule applied? (severity 5 with detectability 4–5 → Band B minimum) |  |
| Urgency triggers (death possibly linked; authority/FDA enquiry; competitor FSCA; media; bench reproduction without mitigation) |  |
| Response time |  |
| Sensitivity check (could a disputed score change the band?) |  |

## Part 4 — Override

| Direction (none / up one band / down one band) |  |
| --- | --- |
| Reason (written; required for any override) |  |
| Chair signature (required for a downward override) and date |  |
| Reviewers |  |

## Re-score log

| Date | Reason for re-score (new data, interim action) | S / P / E / D / V | Total and band | Initials |
| --- | --- | --- | --- | --- |

## Annual calibration record

| Date | Closed signals re-scored blind (n) | Reviewers | Domains with poor agreement | Anchor changes made |
| --- | --- | --- | --- | --- |

Take ten closed signals, have two new reviewers score them blind, and compare. Record the calibration in the PMS plan.

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Glucalyn insulin pump, day 3 (data cut A, Year 2, month 9).

| Domain | Score | Rationale recorded on the worksheet |
| --- | --- | --- |
| Severity | 5 | Severe hypoglycemia; death credible |
| Probability change | 4 | 1 in 2,000 vs 1 in 20,000 pump-years; one window only |
| Exposure | 4 | 24,000 pumps on v3.2; version still being installed |
| Detectability | 5 | Over-delivery silent until hypoglycemia |
| Vulnerability | 4 | Home use by lay users; not a self-rescue-limited population |

Table 13.2 — Glucalyn prioritization at data cut A (illustrative). Total 22 of 25, Band A, no override.

| Field | Entry (Glucalyn, day 3) |
| --- | --- |
| Signal and data cut | Over-delivery on software v3.2; Year 2, month 9 |
| Risk file link | Hazard: over-delivery of insulin; severity scale as in the risk file |
| Scores (S / P / E / D / V) | 5 / 4 / 4 / 5 / 4 |
| Total and band | 22 of 25; Band A (20 to 25) |
| Floor rule applied | Severity 5 with detectability 5: Band B minimum (met by score) |
| Urgency triggers | None on day 3 |
| Response time | Board within 5 working days; action decision within 15 calendar days |
| Override | None. Reason: Band A by score. Chair initials recorded |
| Reviewers | Vigilance lead and safety physician; independent scores agreed |

Table 13.3 — Tool T-11 scoring worksheet, partly filled for the Glucalyn signal (illustrative, company tool).

Sensitivity check recorded: if exposure and vulnerability were each scored 3, the total would be 20, still Band A. Comparison: the Brevant housing-crack series scores 1 + 2 + 4 + 1 + 1 = 9, Band C, despite far higher volume. The FSCA decision came on day 14, within the 15-day Band A time. After the field safety notice on day 21, detectability is re-scored and the re-score recorded; the signal stays in Band A until the patch is verified.