COMPANION TOOLKIT · TOOL T-13

# Event Adjudication Guide and Line-Listing Template

Adjudication rules, line-listing columns with a data dictionary, and a second-reviewer QC checklist

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Turns a retrieval into a set of unique, assessed events. The guide defines each adjudication category for the device problem and the health effect; the line listing records one row per unique event; the QC checklist reconciles the counts before the listing goes to the review board. |
| --- | --- |
| When to use | For every case series built for a signal assessment, including the returned-device flow. Two reviewers adjudicate independently; disagreements are resolved by discussion and recorded. |
| Chapter reference | Chapter 15, Case Series and Returned-Device Investigation — section 15.5 (Tables 15.1–15.3). |

# Part 1 — Adjudication guide

| Axis | Category | Evidence required | Example |
| --- | --- | --- | --- |
| Device problem | Confirmed | Reproduced or observed on the bench | Returned device fails the same way under the documented bench method |
| Device problem | Probable | Clear description, photographs or user-retained evidence; no bench test | Photograph of the failed component retained by the user facility |
| Device problem | Not confirmed | Device tested; no fault found | Returned device passes the bench method; method named in the record |
| Device problem | Unassessable | No device and too little information | Report gives no description beyond "did not work"; device discarded |
| Health effect | Confirmed | Confirmed from clinical information (IMDRF Annex E and F terms) | Operative note or discharge summary documents the outcome |
| Health effect | Reported but unverified | Reporter's statement only | Complaint states a reoperation; no clinical record yet |
| Health effect | None | No health effect described or found | User reports a malfunction noticed before patient contact |

### Questions asked of every event, by the same rules

Device problem — category above, coded in IMDRF Annex A terms.

Health effect — coded separately in IMDRF Annex E and Annex F terms; confirmed from clinical information where available.

Timing — is the harm consistent in time with the device problem?

Alternatives — patient anatomy, tissue quality, technique, concomitant devices.

## Retrieval and deduplication record

| Search strings, sources and date run (so the retrieval can be rerun) |  |
| --- | --- |
| Deduplication rule (shared device identifier, lot, date, facility, narrative) |  |
| Links logged with reasons (reference) |  |

Never delete duplicate records; the master record points to them. Record every return request and every refusal — the reason for non-return is itself data.

## Selection flow

| Stage | Count | Note |
| --- | --- | --- |
| Complaint records retrieved |  |  |
| Unique events after linking |  |  |
| Devices returned |  |  |
| Confirmed on bench |  |  |
| No fault found |  |  |
| No device returned |  |  |

## Data dictionary — line-listing columns

| Column | Content rule | Example entry format |
| --- | --- | --- |
| Event ID and linked records | Master ID plus every linked record ID | Master ID; linked IDs listed |
| Device and lot | Catalog code, lot as labeled | Lot from label; flag if in a suspect lot range |
| Dates and setting | Event date, awareness date, facility type | Operating room, hospital |
| Device problem (IMDRF Annex A) | Coded term plus adjudication category | Term; confirmed / probable / not confirmed / unassessable |
| Health effect (IMDRF Annex E and F) | Coded term, outcome, confirmation source | Term; outcome; source document |
| Return status | Returned / not returned; reason for non-return if known | Not returned — retained by hospital risk department |
| Bench method and result (IMDRF Annex B to D) | Investigation type, finding and conclusion; method named | Method ID; finding; conclusion |
| Alternative explanations | Anatomy, tissue, technique, concomitant devices | Free text, short |
| Reviewer 1 / Reviewer 2 assessment | Independent categories; agreed value after discussion | R1: probable; R2: confirmed; agreed: confirmed |

# Part 2 — Line listing (one row per unique event)

| Event ID / linked records | Device and lot | Dates and setting | Device problem (Annex A) + category | Health effect (Annex E/F), outcome, source | Return status | Bench method and result (Annex B–D) | Alternatives | Reviewer 1 | Reviewer 2 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

# Part 3 — Second-reviewer QC checklist

| Check | Result | Checked by | Date |
| --- | --- | --- | --- |
| Every duplicate link has a stated reason |  |  |  |
| Unique-event total reconciles with the trend count for the same window |  |  |  |
| Returned plus non-returned equals unique events |  |  |  |
| Confirmed plus no fault found equals returned |  |  |  |
| Random sample of rows, plus every row with a serious outcome, compared against source records |  |  |  |
| Every no-fault-found result names the bench method used |  |  |  |
| Every non-returned event has a recorded reason for non-return where one is known |  |  |  |
| Lot proportion paired with distribution share |  |  |  |

Attach the completed checklist to the signal record with the line listing and the flow diagram.

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Suturex circular stapler, trend window Year 3, month 10 to Year 4, month 4.

| Stage | Count | Note |
| --- | --- | --- |
| Complaint records retrieved | 212 | Wide search, window Y3 M10 to Y4 M4 |
| Unique events after linking | 180 | Matches trend total |
| Devices returned | 96 | 53% of events |
| Confirmed misfire on bench | 61 | 44 from lots SX-2310 onward |
| No fault found | 35 | Analyzed, not dismissed |
| No device returned | 84 | Described and adjudicated from records |

Table 15.1 — Suturex selection flow (illustrative, fictional case).

| Column | Content rule | Example entry format |
| --- | --- | --- |
| Event ID and linked records | Master ID plus every linked record ID | Master ID; linked IDs listed |
| Device and lot | Catalog code, cartridge lot as labeled | Lot from label; flag if SX-2310 onward |
| Dates and setting | Event date, awareness date, facility type | Operating room, hospital |
| Device problem (IMDRF Annex A) | Coded term plus adjudication category | Misfire; confirmed / probable / not confirmed / unassessable |
| Health effect (IMDRF Annex E and F) | Coded term, outcome, confirmation source | Anastomotic leak; reoperation; operative note |

Table 15.2 — Tool T-13 line-listing core columns (company tool). The full template adds return status, bench method and result in IMDRF Annex B to D terms, alternative explanations, and each reviewer's assessment.

| Check | Suturex result |
| --- | --- |
| Unique events reconcile with trend total | 180 = 180 |
| Returned plus non-returned equals unique events | 96 + 84 = 180 |
| Confirmed plus no fault found equals returned | 61 + 35 = 96 |
| All serious-outcome rows verified against source | 9 of 9 leak events checked |
| Lot proportion paired with distribution share | Pending distribution extract |

Table 15.3 — Tool T-13 QC reconciliation, filled for Suturex (illustrative, fictional case).