COMPANION TOOLKIT · TOOL T-14

# Root-Cause Investigation and Human-Factors Checklist

Design, manufacturing, use or patient: a fixed structure that makes the evidence for each family of cause visible

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Gives the investigation six parts: problem statement in IMDRF terms; fault tree or fishbone with minimal cut sets; evidence table for design, manufacturing, use and patient; human-factors checklist; attribution statement with its certainty; and the link from each cut set to an action. |
| --- | --- |
| When to use | For any validated signal where the attribution (design, manufacturing, use or patient) decides the action. Store the completed template in the signal record and cross-reference it from the CAPA and the risk management file. |
| Chapter reference | Chapter 16, Root Cause: Design, Manufacturing, Use or Patient — section 16.5 (Tables 16.1 and 16.2). |

| Signal record ID |  | Device / model / component |  |
| --- | --- | --- | --- |
| Data cut |  | Investigation lead |  |

# Part 1 — Problem statement (IMDRF terms)

| Axis | Term (IMDRF) | Description and count |
| --- | --- | --- |
| Device problem (Annex A) |  |  |
| Component (Annex G) |  |  |
| Health effect (Annex E) |  |  |
| Health impact (Annex F) |  |  |

| Pattern and denominator (rates by setting, version, lot) |  |
| --- | --- |

# Part 2 — Fault tree or fishbone

| Top event (precise definition) |  |
| --- | --- |

| ID | Event (basic or intermediate) | Family (design / manufacturing / use / labeling / patient) | Gate it feeds (AND / OR) |
| --- | --- | --- | --- |

| Minimal cut set | Events in the set | Evidence status |
| --- | --- | --- |

With an AND gate at the top, every cut set contains the AND-gate inputs; a control that removes one of them breaks every cut set.

# Part 3 — Evidence table

| Branch | Evidence | Assessment |
| --- | --- | --- |
| Use |  |  |
| Design |  |  |
| Labeling |  |  |
| Manufacturing |  |  |
| Patient |  |  |

# Part 4 — Human-factors checklist (IEC 62366-1 frame)

| Checklist item | Entry |
| --- | --- |
| Task step that failed |  |
| Does the design permit the error? |  |
| Is there an automatic safeguard? |  |
| Representative users tested? |  |
| Error reproduced? |  |
| Use environment considered? |  |
| Labeling salience checked? |  |
| Reportability of use error considered? |  |

# Part 5 — Attribution statement

| Attributed cause(s) and family |  |
| --- | --- |
| Certainty (e.g., established / probable / possible) and reasoning |  |
| Families considered and not supported, with evidence |  |

# Part 6 — Cut set to action

| Minimal cut set | Action that breaks it | Interim or permanent | Effectiveness check | CAPA / risk file reference |
| --- | --- | --- | --- | --- |

### Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Investigation lead |  |  |  |
| Human-factors specialist |  |  |  |
| Clinical reviewer |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Orvana infusion pump with the C-2 cassette, Year 4 (45 free-flow complaints; home 7.0 vs hospital 1.6 per 1,000 pumps).

| ID | Event | Family / gate |
| --- | --- | --- |
| Top | Free-flow event: uncontrolled gravity infusion from an Orvana pump with the C-2 cassette | AND gate G1 (B1, B2, E1) |
| B1 | C-2 cassette has no automatic anti-free-flow clamp | Design |
| B2 | Pump permits cassette removal while the flow path is open (no clamp interlock) | Design |
| E1 | Flow path open when the cassette is removed | OR gate G2 (E2, B6) |
| E2 | Roller clamp not closed before removal | Use; OR gate G3 (B3, B4, B5) |
| B3 | Clamping step omitted during an interrupted task (bag change, alarm handling, end of infusion) | Use |
| B4 | Clamping step not trained, or not retained, by a lay caregiver | Use |
| B5 | Clamping cue in the IFU or on the device not seen at the point of use | Labeling (a design family) |
| B6 | Roller clamp closed but does not occlude the tubing | Manufacturing |

Fault tree from Chapter 16, Step 2. Minimal cut sets: {B1, B2, B3}, {B1, B2, B4}, {B1, B2, B5}, {B1, B2, B6}.

| Branch | Evidence | Assessment |
| --- | --- | --- |
| Use (B3, B4) | Home rate 7.0 vs hospital 1.6 per 1,000 pumps; 7 of 9 consequences at home; simulated use 4 of 15 lay caregivers vs 1 of 15 nurses | Supported; stronger with lay users |
| Design (B1, B2) | Present in every C-2 unit; IFU instructs clamping but design does not enforce it; error reproduced in simulated use | Supported; necessary in every cut set |
| Labeling (B5) | Clamping step is in the IFU; salience at point of use not yet tested | Possible; test in revised training |
| Manufacturing (B6) | Same pump in both settings, yet rates differ by setting; confirm with returned-cassette testing and lot genealogy | Not supported by the pattern; close with records |
| Patient | Does not cause free flow; affects severity of harm | Record for the risk file, not the cause |

Table 16.1 — Orvana root-cause evidence table (illustrative, fictional case).

| Checklist item | Orvana entry |
| --- | --- |
| Task step that failed | Closing the roller clamp before cassette removal |
| Does the design permit the error? | Yes: removal allowed with the flow path open (B2) |
| Is there an automatic safeguard? | No: C-2 has no anti-free-flow clamp (B1) |
| Representative users tested? | Lay caregivers and nurses, 15 each |
| Error reproduced? | 4 of 15 lay caregivers; 1 of 15 nurses |
| Use environment considered? | Home: interruptions, night-time bag changes, no second person |
| Labeling salience checked? | Pending; include in revised home training |
| Reportability of use error considered? | Yes, per jurisdiction (GB rule on use error noted) |

Table 16.2 — Tool T-14 human-factors checklist, partly filled for Orvana (illustrative, company tool).

Part 6 for Orvana: every cut set is broken by the C-3 cassette (automatic anti-free-flow clamp removes B1); retraining acts on B4 and partly on B3; the interim IFU and home-training update and home-kit caution label reduce three of the four cut sets until conversion is complete. Effectiveness check: home free-flow complaints below 2.0 per 1,000 pumps per year (annualized from each quarter's count) for two consecutive quarters after C-3 conversion.