COMPANION TOOLKIT · TOOL T-15

# Totality-of-Evidence and Device-Causality Template

Part A: evidence table · Part B: structured causality argument · Part C: four-part conclusion

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Turns integrated assessment into a record: one row per evidence source with its weight for the question (Part A); a device-adapted causality argument (Part B); and a conclusion that states what the evidence supports, what remains uncertain, the effect on the risk management file and the action that follows (Part C). |
| --- | --- |
| When to use | When the parts of an assessment (question, case series, returned-device findings, root cause) are complete. Complete the template before the review board meets. |
| Chapter reference | Chapter 17, Integrated Assessment and Causality — section 17.5 (Tables 17.1 and 17.2). |

| Signal record ID |  | Device / constituent |  |
| --- | --- | --- | --- |
| Data cut |  | Assessment question (T-12 reference) |  |

# Part A — Evidence table

Weight belongs to the pairing of a source and a question. For each source ask: does it measure the outcome the question defines? Does it cover the population, version, lot or technique in question? Does it have a denominator and a fair comparator?

| Source | Finding | Strength and limitation | Direction | Weight |
| --- | --- | --- | --- | --- |

### Discordance log — explain each disagreement before concluding

| Sources that disagree | Explanation (outcome definition, population, technique, time window, underreporting) |
| --- | --- |

# Part B — Device-causality argument

| Consideration | Evidence | Effect on judgment (supports / weakens / neutral) | Gap to close |
| --- | --- | --- | --- |
| Temporality |  |  |  |
| Consistency |  |  |  |
| Dose-like relationship (time in use, cycles, size, amount of substance) |  |  |  |
| Removal or replacement response |  |  |  |
| Bench reproduction |  |  |  |
| Plausibility (mechanism) |  |  |  |
| Evidence from similar devices |  |  |  |

When a row says "neutral", write down why. A blank row and a neutral row mean different things to an auditor. No single consideration is required, and none is enough alone.

# Part C — Conclusion (four parts, in order)

| 1. What the evidence supports (confirmed / refuted / indeterminate; graded causal judgment) |  |
| --- | --- |
| 2. What remains uncertain, and the evidence that would reduce it |  |
| 3. Effect on the risk management file (new hazardous situation, changed probability or severity) — RMF reference |  |
| 4. Action that follows — linked action numbers |  |

### Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Signal owner |  |  |  |
| Clinical lead |  |  |  |
| Engineering lead |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Veltrix PEEK cage pre-loaded with OIP-4 (combination product), data cut Year 3, month 6.

| Source | Finding | Strength and limitation | Direction | Weight |
| --- | --- | --- | --- | --- |
| Complaints (19) | Onset 2–14 days; 12 off-label; 4 ectopic bone | Timing and imaging informative; no denominator by approach | Supports | Moderate |
| PMCF (7 of 210) | 3.3% (1.6–6.7) in anterior use | Active case finding; small; external comparator | Supports, on-label too | High for on-label incidence |
| Literature | About 1% without protein | Different populations and definitions | Sets comparator | Moderate |
| Spine registry | Weak signal | Outcome is reoperation; misses non-surgical cases | Neutral | Low |
| Engineering | No cage failure | Excludes a mechanical cause | Relocates the cause | High for that point |
| Mechanism | Protein escape, root exposure | Fits timing and imaging | Supports | Moderate |

Table 17.1 — Veltrix evidence table, data cut Year 3, month 6 (Part A, illustrative).

| Consideration | Veltrix evidence | Effect on judgment | Gap to close |
| --- | --- | --- | --- |
| Temporality | Onset median 5 days (2–14) | Supports | None |
| Consistency | Complaints, PMCF and mechanism agree; registry explained | Supports | Standard definition in PMCF |
| Dose-like relationship | Off-label approaches, which expose the root, dominate reports | Supports, not quantified | Approach-specific denominator |
| Removal or replacement | Not applicable to a fused cage | Neutral | None feasible |
| Reproduction and plausibility | No cage failure; ectopic bone in 4 of 19; inflammatory mechanism | Supports a constituent cause | Imaging review of PMCF cases |

Table 17.2 — T-15 Part B, device-causality argument for Veltrix (illustrative).

| Part C | Veltrix entry |
| --- | --- |
| 1. What the evidence supports | Signal confirmed. Probable causal association between OIP-4 delivered with Veltrix and postoperative radiculitis; strongest with off-label posterior or transforaminal use; some excess also with anterior use. |
| 2. What remains uncertain | The size of the on-label excess and the true rate by approach. |
| 3. Effect on the risk management file | New hazardous situation (protein reaching an exposed nerve root); re-estimated probability for on-label use. |
| 4. Action that follows | IFU warning and contraindication for posterior and transforaminal approaches; surgeon communication; PMCF amendment adding a standard radiculitis definition; report, through the notified body, to the medicines authority consulted on the ancillary substance (MDR Article 1(8), Annex IX section 5.2). |

Part C summarized from Chapter 17, Step 8. Drug-constituent events are coded in MedDRA and the device aspects in IMDRF terms.