COMPANION TOOLKIT · TOOL T-16

# Risk Management File Update Worksheet

Before the signal, interim, and after verification — with the evidence for each change

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Records the three updates to one risk management file entry in one place: the value before the signal, the interim value and the post-verification value, with the evidence behind each change. The header carries the signal record number so the link runs both ways. |
| --- | --- |
| When to use | When a signal changes a hazardous situation, a severity, or a probability of harm (ISO 14971 clause 10 production and post-production review). Revisit after interim controls and again after verification of the permanent control. |
| Chapter reference | Chapter 18, From Signal to Risk Management File — section 18.5 (Tables 18.1–18.3). Companion web tool: https://abdu94-hash.github.io/device-signal-tool/ |

| Signal record number |  | Risk management file / hazard ID |  |
| --- | --- | --- | --- |
| Device / version |  | Prepared by / date |  |

# Part 1 — Post-market probability estimate

☐  Use the exposure of the affected population (for example, device-years on the affected version, not the whole fleet).

☐  State the uncertainty. A small count gives a wide interval; the risk evaluation should hold across the whole interval, not only at its midpoint.

☐  Write down what the estimate cannot see (unreported events, short follow-up).

| Quantity | Value | Source / calculation |
| --- | --- | --- |
| Events with the harm of interest (n) |  |  |
| Exposure of the affected population |  |  |
| Estimate = n ÷ exposure ("1 in …") |  |  |
| Interval (method) |  |  |
| Risk file estimate before the signal |  |  |
| Known gaps in the estimate |  |  |

# Part 2 — Update worksheet

| Risk file field | Before signal | Interim | After verification (target / verified) | Evidence for the change |
| --- | --- | --- | --- | --- |
| Sequence of events |  |  |  |  |
| Severity |  |  |  |  |
| Probability of harm |  |  |  |  |
| Risk region |  |  |  |  |
| Controls and evidence |  |  |  |  |

Interim controls that rely on information (for example, a field safety notice) are recorded but, as in the book's example, may be given no risk-reduction credit. Record the reasoning when severity is unchanged.

# Part 3 — Review questions

| Movement within a region: is further reduction now practicable, and does the benefit still outweigh the risk? |  |
| --- | --- |
| State of the art: do comparable devices now include a protective measure this device lacks? |  |
| Does the overall residual risk and benefit–risk conclusion still hold? |  |

### Which documents inherit the change?

| Document | Change needed (Y/N) | Owner | Due / done |
| --- | --- | --- | --- |
| Clinical evaluation |  |  |  |
| PSUR |  |  |  |
| IFU |  |  |  |
| Field safety notice |  |  |  |
| Other (SSCP, PMCF plan, PMS plan) |  |  |  |

Use Tool T-22 to reconcile the documents.

### Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Risk management lead |  |  |  |
| Safety physician |  |  |  |
| Signal owner |  |  |  |

# Reference — illustrative company risk matrix

A company example, not a regulatory requirement. ISO 14971 leaves the acceptability criteria to the manufacturer's policy, which must reflect the generally acknowledged state of the art. "ALARP region" is the company's label for risks acceptable only after showing that further reduction is not practicable and that the benefit outweighs the residual risk.

| Probability band | Occurrences per pump-year | Acceptable at | ALARP region at | Unacceptable at |
| --- | --- | --- | --- | --- |
| P5 Frequent | 1 in 1,000 or more | None | S1–S2 | S3–S5 |
| P4 Probable | 1 in 10,000 to less than 1 in 1,000 | S1 | S2–S3 | S4–S5 |
| P3 Occasional | 1 in 100,000 to less than 1 in 10,000 | S1–S2 | S3–S4 | S5 |
| P2 Remote | 1 in 1,000,000 to less than 1 in 100,000 | S1–S3 | S4–S5 | None |
| P1 Improbable | Less than 1 in 1,000,000 | S1–S4 | S5 | None |

Table 18.1 — Illustrative company 5×5 risk matrix: probability bands and risk regions.

| Level | Label | Definition (company example) |
| --- | --- | --- |
| S1 | Negligible | Inconvenience or temporary discomfort; no intervention |
| S2 | Minor | Temporary injury managed by the user or patient |
| S3 | Serious | Injury needing professional medical intervention; reversible |
| S4 | Critical | Life-threatening injury, third-party rescue, hospitalization or permanent impairment |
| S5 | Catastrophic | Death |

Table 18.2 — Illustrative company severity levels.

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Glucalyn insulin pump, over-delivery on software v3.2 (data cut A, Year 2, month 9).

| Quantity | Glucalyn value |
| --- | --- |
| Events with severe harm | 6, all on v3.2 (2 hospitalized; no deaths) |
| Exposure of the affected population | 12,000 pump-years on v3.2 (months 4–9) |
| Estimate | 6 ÷ 12,000 = 1 in 2,000 pump-years |
| Interval | About 1 in 5,450 to 1 in 920 |
| Risk file estimate before the signal | 1 in 20,000 pump-years |

| Risk file field | Before signal | Interim (day 14–21) | After patch (target) |
| --- | --- | --- | --- |
| Sequence of events | Software timing fault (general) | v3.2: alarm cleared with bolus pending; bolus delivered twice | Same sequence; fault removed in v3.2.1 |
| Severity | S4 | S4 (reasoning recorded) | S4 |
| Probability of harm | 1 in 20,000 pump-years (P3) | 1 in 2,000 (P4); interval about 1 in 5,450 to 1 in 920 | 1 in 200,000 or better (P2), verification target |
| Risk region | ALARP region | Unacceptable | ALARP region, once verified |
| Controls and evidence | Software verification of v3.2 | FSCA; field safety notice with interim instruction (no credit) | Timing fix plus independent volume check; verification report; field monitoring |

Table 18.3 — T-16 risk file update worksheet for Glucalyn over-delivery (illustrative).