COMPANION TOOLKIT · TOOL T-17

# Action-Selection Matrix and FSCA Decision Record

Part A: grade the factors and choose actions across the ladder · Part B: document one decision so an auditor can follow it

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Part A grades evidence strength, severity, exposure, preventability and alternatives, and maps the pattern to the rungs of the action ladder (a company example). Part B documents one decision — scope, rejected options, effectiveness criteria, owners, dates, notifications and record links. |
| --- | --- |
| When to use | When an assessed signal needs an action decision, including interim containment before the root cause is known. |
| Chapter reference | Chapter 19, Selecting the Action — section 19.5 (Tables 19.1 and 19.2). |

| Signal record ID |  | Device / lots / versions |  |
| --- | --- | --- | --- |
| Data cut |  | Prepared by / date |  |

# Part A — Action-selection matrix

| Factor | Lower | Higher | Grade and evidence |
| --- | --- | --- | --- |
| Evidence strength | Alert or unconfirmed pattern; cause unknown | Confirmed signal; cause identified; bench reproduction |  |
| Severity | Reversible, minor harm | Serious, irreversible or life-threatening harm |  |
| Exposure | Few units; affected stock consumed | Wide distribution; affected configuration still in use or shipping |  |
| Preventability by users | User can detect and avoid the failure | User cannot detect or avoid it; instructions add little |  |
| Alternatives and cost of the action | No alternative; action itself harms patients | Unaffected lots, versions or devices available |  |

### Mapping to the ladder (company example)

A signal enters at the rung the evidence supports; it does not climb one rung at a time, and actions from several rungs often run together.

| Rung | Action | Typical pattern (company example) | Selected? (Y/N) |
| --- | --- | --- | --- |
| 1 | Routine monitoring | Weak evidence, low severity |  |
| 2 | Enhanced monitoring and targeted evidence | Evidence not yet strong enough for a field action; give it a question, an end date and a decision |  |
| 3 | Information updates | Cause lies in use and the design cannot reasonably enforce safe use |  |
| 4 | Field safety notice with interim instructions | Users can reduce harm while a permanent fix is prepared |  |
| 5 | Correction on site or by software | Defined fix can be applied to distributed devices |  |
| 6 | Removal or recall of a defined scope | Confirmed defect in defined lots or versions; users cannot prevent the failure; alternatives exist |  |
| 7 | Market withdrawal or suspension | Unacceptable risk across the product with no adequate fix |  |

### Actions by class

| Class | Action | What it prevents | Scope |
| --- | --- | --- | --- |
| Evidence-generating |  |  |  |
| Field |  |  |  |
| Design / process |  |  |  |
| Information and training |  |  |  |
| Risk file / monitoring |  |  |  |

Base scope on the cause, not on the complaints: a lot with no complaints yet may still carry the defect. For implants, scope also covers patients who already have the device.

# Part B — FSCA decision record

| Signal and evidence |  |
| --- | --- |
| Harm and exposure |  |
| Decision and scope |  |
| Interim action, permanent action and handover between them |  |
| Patients already treated (advisory needed? reasoning) |  |
| Options rejected, with reasons |  |
| Effectiveness — execution criterion (units found, retrieved or corrected; consignee acknowledgment) |  |
| Effectiveness — outcome criterion (event rate set in advance) and review date |  |
| Links: signal record, trend report, risk file revision, CAPA number |  |

### Regulatory notifications made before the action starts

| Authority / jurisdiction | Notification | Date sent | Logged by |
| --- | --- | --- | --- |

MDR: an FSCA is reported before it is undertaken, except in urgent cases, and communicated through a field safety notice. Great Britain: the proposed notice goes to MHRA before customers. US: a correction or removal to reduce a risk to health is reported within 10 working days of starting it (21 CFR Part 806); recalls are classified under Part 7.

### Decision

| Decision-maker (name, role) | Signature | Date | Review board minute reference |
| --- | --- | --- | --- |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Suturex circular stapler, data cut Year 4, month 4.

| Factor | Suturex grading (Chapter 19, Step 2) |
| --- | --- |
| Evidence strength | Trend significant under the PMS plan rule; 61 confirmed misfires, 44 from lots SX-2310 onward; cause is the anvil-spring supplier change — confirmed, supplier-related |
| Severity | 9 anastomotic leaks needing reoperation: serious, not reversible without surgery |
| Exposure | 50,000 units sold each month |
| Preventability | A surgeon cannot detect a weak spring before firing, so instructions cannot prevent the failure |
| Alternatives | Cartridges from earlier lots and from the corrected supply perform at the labeled rate |

| Class | Action | What it prevents | Scope |
| --- | --- | --- | --- |
| Field | FSCA with recall and quarantine | Use of defective cartridges already distributed | Lots SX-2310 to SX-2404 |
| Field / information | Field safety notice to surgical teams | Use of stock before retrieval; missed leak | All consignees of those lots |
| Design / process | Reverse the anvil-spring supplier change | New defective lots | All future production |
| Process | CAPA with effectiveness criterion | Recurrence through supplier control | Supplier qualification and incoming inspection |
| Risk file / monitoring | Raise misfire severity to include leak; continue p-chart | Understated risk; missed recurrence | Whole product line |

Table 19.1 — Suturex actions by class, data cut Year 4, month 4 (illustrative).

| Field | Entry |
| --- | --- |
| Signal and evidence | Confirmed, supplier-related; 44 of 61 confirmed misfires from SX-2310 onward; Article 88 trend report filed |
| Harm and exposure | 9 anastomotic leaks with reoperation among 180 events; 50,000 units a month |
| Decision and scope | FSCA and recall of lots SX-2310 to SX-2404; field safety notice to surgical teams; supplier change reversed |
| Options rejected, with reasons | Reminder letter alone (cause not in user control); reversal without field action (stock stays in use); withdrawal (other lots at labeled rate) |
| Effectiveness and review | Rate below 0.48 per 1,000 for three consecutive months; full lot reconciliation; review after three months of reversed-supply data |

Table 19.2 — T-17 FSCA decision record for Suturex, partly completed (illustrative).

Patients already treated: single-use device, so no device remains in patients; the notice reminds teams of the leak finding and no separate patient advisory is issued, with the reasoning recorded. The 0.48 per 1,000 criterion means fewer than 24 misfires a month at 50,000 units — stricter than returning inside the 2σ limit of 28.9. The record also needs the decision-maker, the date, the notifications made before the action starts, and the links to the CAPA and the risk file revision.