COMPANION TOOLKIT · TOOL T-18

# Device Benefit–Risk Worksheet

Effects, patient groups, subgroups and value judgments — in a form the board can sign

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Part A lists favorable and unfavorable effects with frequency, comparator, uncertainty and burden. Part B separates implanted patients from future patients. Part C covers subgroups. Part D records the value judgments and their owners. An optional quantitative appendix holds the per-10,000 conversion, the uncertainty bracket and the break-even weight. |
| --- | --- |
| When to use | When a confirmed or likely signal changes the balance for a device or a configuration; when new data arrive, reopen the worksheet rather than rewrite it — the new version cites the old one and shows which judgments changed. |
| Chapter reference | Chapter 20, Benefit–Risk for Devices — section 20.5 (Tables 20.1 and 20.2). |

| Device / configuration |  | Signal record ID |  |
| --- | --- | --- | --- |
| Data cut |  | Worksheet version (supersedes) |  |

# Part A — Effects

| Effect (favorable / unfavorable) | Device | Comparators | Per 10,000 | Uncertainty, burden and notes |
| --- | --- | --- | --- | --- |

# Parts B and C — Patient groups and subgroups

| Group | Balance on current evidence | Decision | Evidence still needed |
| --- | --- | --- | --- |
| Future patients — affected configuration |  |  |  |
| Future patients — other configurations |  |  |  |
| Implanted / already treated patients |  |  |  |
| Subgroup: |  |  |  |
| Subgroup: |  |  |  |

# Part D — Value judgments

| Judgment | Owner | Signature | Date |
| --- | --- | --- | --- |

# Appendix — Quantitative (optional)

| Step | Formula | Value |
| --- | --- | --- |
| Unfavorable events added per 10,000 | (device rate − comparator rate) × 10,000 |  |
| Favorable events prevented per 10,000 | (comparator rate − device rate) × 10,000 |  |
| Rough uncertainty bracket (not a formal interval for the difference) | Low: device lower limit − comparator upper limit; High: device upper limit − comparator lower limit |  |
| Break-even weight | Events added ÷ events prevented |  |
| Reverse check | Events prevented ÷ events added |  |

Do not net the two counts into one number; show them side by side, and decide by group.

### Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Review board chair |  |  |  |
| Clinical lead |  |  |  |
| Signal owner |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Tessora hip, Year 6 registry report.

| Effect | Tessora | Comparators | Per 10,000 | Uncertainty and notes |
| --- | --- | --- | --- | --- |
| Dislocation at 5 years (favorable) | 0.8% | 1.5% | 70 prevented | No interval given; overlaps with revision count |
| Revision at 5 years (unfavorable) | 4.1% (3.3–5.1) | 2.4% (2.2–2.6) | 170 added | Rough bracket 70–290; O/E 1.85 (1.36–2.45) |
| Liner dissociation (unfavorable) | 21 of 48 revisions | Not reported | Not estimated | 17 of 21 in size 52–54, made after Year 3, month 4 |
| Burden of each event | Revision: major reoperation | Dislocation: often closed reduction | Not netted | Break-even weight 2.4 revisions per dislocation |

Table 20.1 — Tessora structured benefit–risk summary (Part A, illustrative).

| Group | Balance on current evidence | Decision | Evidence still needed |
| --- | --- | --- | --- |
| Future patients, size 52–54, affected period | Unfavorable | Quarantine under FSCA; CAPA on tolerance | Confirm the CAPA restores performance |
| Future patients, other configurations | Uncertain; not shown to be unfavorable | Continue; review O/E at each annual report | Revision by size and period |
| Implanted patients, any size | Revision burden outweighs a prophylactic benefit | No prophylactic revision; radiographs at next visit | Registry follow-up of affected sizes |
| Older, low-demand patients | Benefit and burden both larger | Surgeon judgment with the patient | Age-specific registry data |

Table 20.2 — T-18 Parts B and C for Tessora, partly completed (illustrative).

| Part D judgment (Tessora) | Owner |
| --- | --- |
| A revision weighs more than a dislocation. | Board chair (signed) |
| The aggregate figure does not decide the fate of unaffected sizes. | Board chair (signed) |
| Patients already implanted should not be exposed to surgery for a risk most of them will never meet. | Board chair (signed) |

| Step | Tessora calculation |
| --- | --- |
| Revisions added per 10,000 | (4.1% − 2.4%) × 10,000 = 170 |
| Dislocations prevented per 10,000 | (1.5% − 0.8%) × 10,000 = 70 |
| Rough bracket | 3.3 − 2.6 = 0.7 points (70) to 5.1 − 2.2 = 2.9 points (290) |
| Break-even weight | 170 ÷ 70 = 2.4 revisions per dislocation |
| Reverse check | Balance favors Tessora only if a revision counted for less than 70 ÷ 170 = 0.41 of a dislocation |

Quantitative appendix, Chapter 20, Steps 1–4.