COMPANION TOOLKIT · TOOL T-19

# Field Safety Notice Checklist

Content, order, language, regulatory sequence and follow-up — checked on every draft from first version to issue

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Checks a field safety notice draft by draft, so that readers can tell within the first lines whether they are affected and what to do. Includes an audience plan, the recommended section order and a plain-language rewrite exercise. |
| --- | --- |
| When to use | On every draft of a field safety notice or advisory, from first version to issue. File the completed checklist with the FSCA decision record (Tool T-17). |
| Chapter reference | Chapter 21, Communicating Device Signals — section 21.5 (Tables 21.1–21.3). |

| Device / versions / lots |  | FSCA reference |  |
| --- | --- | --- | --- |
| Draft versions checked |  | Checklist owner |  |

# Part 1 — Audiences and routes

| Audience | What they need first | Route | Owner | Planned date |
| --- | --- | --- | --- | --- |
| Competent authorities / notified body (EU) |  |  |  |  |
| MHRA (Great Britain) |  |  |  |  |
| FDA (United States) |  |  |  |  |
| Clinicians / prescribers |  |  |  |  |
| Users, patients and caregivers |  |  |  |  |
| Distributors and service providers |  |  |  |  |

# Part 2 — Draft checklist

| Check | Draft 1 | Draft 2 | Final | Owner |
| --- | --- | --- | --- | --- |
| Title says "Urgent Field Safety Notice" and names the device |  |  |  |  |
| Affected devices identified by the attribute that carries the fault, with how to check |  |  |  |  |
| Unaffected users told whether they need to act |  |  |  |  |
| Required action in the first lines, as one imperative sentence |  |  |  |  |
| Interim action separated from the permanent correction |  |  |  |  |
| No release date promised for an unverified correction |  |  |  |  |
| Risk stated in plain words, with the symptoms that call for action |  |  |  |  |
| Response form and pass-on instructions included |  |  |  |  |
| Contacts, and a statement that the authorities have been informed |  |  |  |  |
| Proposed notice sent to MHRA before customers; FSCA reported in the EU before the action starts |  |  |  |  |
| Effectiveness check defined in the FSCA file |  |  |  |  |

# Part 3 — Recommended section order

| Position | Section | Present in draft? (Y/N) |
| --- | --- | --- |
| 1 | "Urgent Field Safety Notice", device, version or lot, one-line risk |  |
| 2 | Is your device affected? (how to check; who is not affected) |  |
| 3 | What to do now: the interim instruction |  |
| 4 | What happens next: the permanent correction, and how users will be contacted |  |
| 5 | The risk in plain words, and the symptoms that call for action |  |
| 6 | What went wrong, in two sentences |  |
| 7 | Response form and instructions to pass the notice on |  |
| 8 | Contacts, and a statement that the authorities have been informed |  |

# Part 4 — Plain-language rewrite exercise

Take one sentence from the technical draft and rewrite it twice, in no more than 60 words each. Test both versions with someone outside the project. If they cannot say what to do after reading it, rewrite it again.

| Original sentence |  |
| --- | --- |
| Clinician version (≤ 60 words) |  |
| Patient version (≤ 60 words) |  |
| Tested with (role) — could they say what to do? |  |

# Part 5 — Follow-up

| Measure | Target | Result | Date |
| --- | --- | --- | --- |
| Consignees reached / acknowledged |  |  |  |
| Response forms returned |  |  |  |
| Non-responders followed up in writing |  |  |  |
| Units corrected / retrieved |  |  |  |

### Approval to issue

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Vigilance lead |  |  |  |
| Clinical lead |  |  |  |
| Regulatory affairs |  |  |  |
| Quality lead |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Glucalyn insulin pump, first draft and revised notice.

| Check | First draft | Revised notice | Owner |
| --- | --- | --- | --- |
| Title says "Urgent Field Safety Notice" and names the device | No | Yes | Vigilance lead |
| Affected devices identified by the attribute that carries the fault, with how to check | No (catalog number) | Yes (software v3.2; information screen) | Vigilance lead |
| Unaffected users told whether they need to act | No | Yes (v3.1) | Vigilance lead |
| Required action in the first lines, as one imperative sentence | No (page two) | Yes | Clinical lead |
| Interim action separated from the permanent correction | No | Yes | Clinical lead |
| No release date promised for an unverified correction | Not applicable | Yes | Engineering lead |
| Response form and pass-on instructions included | No | Yes | Quality lead |
| Proposed notice sent to MHRA before customers; FSCA reported in the EU before the action starts | Not started | Planned with dates | Regulatory affairs |
| Effectiveness check defined in the FSCA file | No | Yes | Quality lead |

Table 21.3 — Tool T-19 extract, completed for the Glucalyn drafts.

| Position | First draft | Revised notice |
| --- | --- | --- |
| 1 | "Important Product Information" and company letterhead | "Urgent Field Safety Notice", device, software version, one-line risk |
| 2 | Software architecture and v3.2 release history | Is your pump affected? (v3.2 only; how to check; v3.1 not affected by this problem) |
| 3 | Engineering description of the timing fault | What to do now: the interim instruction |
| 4 | Investigation summary and complaint statistics | What happens next: update v3.2.1, and how users will be contacted |
| 5 | Affected products by catalog number | The risk in plain words, and the symptoms that call for action |
| 6 | Risk to patients, after a reassuring sentence | What went wrong, in two sentences |
| 7 | Actions to be taken | Response form and instructions to pass the notice on |
| 8 | Contact details | Contacts, and a statement that the authorities have been informed |

Table 21.2 — Glucalyn field safety notice: section order before and after revision. Interim instruction: after clearing any occlusion alarm, check the delivery history before giving another bolus.

| Audience | What they need first | Route | Owner |
| --- | --- | --- | --- |
| EU competent authorities; notified body | FSCA report and draft notice before the action starts | Vigilance reporting route; notified body contact | Regulatory affairs |
| MHRA (Great Britain) | Proposed notice before customers receive it, with UDI where available | MHRA reporting route | Regulatory affairs |
| FDA (United States) | Report of the correction within 10 working days of starting it | 21 CFR Part 806 report | Regulatory affairs |
| Diabetes clinics and prescribers | Affected versions, the interim instruction, the clinical risk in numbers | Clinician letter with the notice | Clinical lead |
| Pump users and caregivers | Is my pump affected; what do I do now | Notice in plain language | Vigilance lead |
| Distributors and home-supply providers | Which customers to reach; how to confirm onward delivery | Notice plus distributor instructions | Quality and supply chain |

Table 21.1 — Audiences for the Glucalyn field safety notice.