COMPANION TOOLKIT · TOOL T-20

# Post-Market Review Board Charter, Agenda, Decision Request and Minutes

A board that exists to decide: charter, 60-minute agenda, decision request cover sheet, decision log and minutes template

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Gives the post-market review board a written charter, an agenda built around decisions, a one-page decision request for every item, a decision log that carries open actions across meetings, and a minutes template that records decisions rather than discussion. |
| --- | --- |
| When to use | When setting up or revising the board; before every routine or urgent session. Adapt quorum, decision rights and urgent-path triggers, then test them against a past case. |
| Chapter reference | Chapter 22, Post-Market Review Board Meetings — section 22.5 (Tables 22.1–22.4). |

# Part 1 — Charter

Company-designed example text in grey cells; replace with your organization's choices.

| Charter element | Example text (adapt) | Your organization |
| --- | --- | --- |
| Purpose | Turn evidence into decisions that one function cannot take alone: validation, trend-report criteria, assessment completeness, recommended action, reportability confirmation, escalation. |  |
| Quality-system basis | MDR Article 10(9) places PMS and vigilance inside the QMS; ISO 13485:2016 clauses 8.2.1 to 8.2.3. Charter, agenda, decision requests and minutes are quality records. |  |
| Standing members | Vigilance (chair or secretary); clinical or medical; R&D and engineering; quality and CAPA; regulatory affairs; risk management; manufacturing. Add supplier quality, human factors or software specialists as needed. |  |
| Quorum | Chair, a clinical member, a regulatory member, a quality member, and at least one of engineering or risk management. |  |
| Conflicts of interest | Declared at the start of each meeting and recorded in the minutes. |  |
| Chair | Owns the agenda and the clock; accepts an item only with a decision request; orders items by urgency; when time runs out sets a named owner, a date and a route back to the board. |  |
| Secretary | Builds the pack; checks each decision request for a clear question and data cut; drafts decision lines live; carries open actions from the decision log. |  |
| Decision rights | Decides validation, trend-report filing, assessment scope and closure. Recommends FSCAs and recalls for approval by a named executive. Escalates to management review any decision with major supply or resource impact; escalation carries a deadline and does not pause safety work. |  |
| Dissent | Recorded, not smoothed over: member, reason, and why the board proceeded. |  |
| Cadence | Monthly for most portfolios. |  |
| Urgent-path triggers | New or more severe harm; lot cluster with injury; 3σ breach; suspected serious public health threat. |  |
| Urgent convening | Who convenes a quorate subset, and how fast: [to be set by the company]. |  |

### Charter approval

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Board chair |  |  |  |
| Head of quality |  |  |  |
| Named executive (FSCA approval) |  |  |  |

# Part 2 — 60-minute agenda

| Meeting date |  | Routine / urgent |  |
| --- | --- | --- | --- |
| Chair |  | Secretary |  |

| Block | Content | Minutes | Items and presenters |
| --- | --- | --- | --- |
| 1. Opening | Quorum, conflicts, open actions from the decision log | 5 |  |
| 2. Priority decision | Items on the urgent path or with a regulatory clock | 20 |  |
| 3. Other decision items | Validations, assessment scopes, closures, each with a decision request | 15 |  |
| 4. Reportability and escalation | Reportability confirmed for each decision; items escalated with deadlines | 10 |  |
| 5. Consent block | Trend dashboards and reviews with no change, taken as read | 5 |  |
| 6. Read-back | Decisions, owners and due dates read aloud and agreed | 5 |  |
| Total |  | 60 |  |

Charts that need no decision go to the consent block; any member can pull an item into discussion.

# Part 3 — Decision request cover sheet (one page per item)

| Item and product |  |
| --- | --- |
| Data cut |  |
| Decision required (yes/no, or a choice between named options) |  |
| Evidence (appendix IDs) |  |
| Options and consequences |  |
| Recommendation |  |
| Reportability |  |
| Deadline |  |
| Escalation needed |  |
| Prepared by / date |  |

If the author cannot write the question, the item is not ready for the board. Charts go in an appendix.

# Part 4 — Decision log

| No. | Decision | Rationale | Owner | Due | Status |
| --- | --- | --- | --- | --- | --- |

# Part 5 — Minutes template

| Date and time |  | Routine / urgent |  |
| --- | --- | --- | --- |
| Chair |  | Secretary |  |
| Members present |  | Quorum met (Y/N) |  |
| Conflicts declared |  | Open actions reviewed (Y/N) |  |

### Per item (repeat)

| Item and decision request ID |  |
| --- | --- |
| Evidence considered (document IDs and data cut) |  |
| Decision |  |
| Rationale |  |
| Owners and due dates |  |
| Reportability position |  |
| Dissent (member, reason, why the board proceeded) |  |

☐  Read-back completed: each decision read aloud; each owner confirmed action and date.

☐  No item recorded only as "discussed" or "noted" — these are not decisions.

| Minutes approved by (chair) | Signature | Date |
| --- | --- | --- |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Suturex circular stapler, urgent session at Year 4, month 4.

| Field | Entry |
| --- | --- |
| Item and product | Suturex circular stapler: misfire complaints |
| Data cut | Year 4, month 4; window Year 3, month 10 to Year 4, month 4 |
| Decision required | (1) Is the PMS plan trend-report criterion met? (2) Is this a validated signal needing assessment and a risk file review? |
| Evidence (appendix IDs) | Monthly counts and limits (Table 22.2); list of the 9 leak events; supplier change record for lots from SX-2310 |
| Options | Q1: file a trend report, or record why the criterion is not met. Q2: validate and assess, or keep monitoring |
| Recommendation | File the trend report; validate as a signal; stratify the assessment by lot |
| Reportability | Article 88 trend report; individual reportability of the 9 leak events to be confirmed |
| Deadline | Decision today (urgent path) |
| Escalation needed | Yes, if an FSCA is recommended |
| Prepared by / date | Vigilance lead / Year 4, month 5 |

Table 22.4 — Tool T-20 decision request cover sheet, completed for Suturex.

| Month | Misfire complaints | Rate per 1,000 units | Beyond 2σ (28.9)? | Beyond 3σ (33.4)? |
| --- | --- | --- | --- | --- |
| Y3 M10 | 19 | 0.38 | No | No |
| Y3 M11 | 23 | 0.46 | No | No |
| Y3 M12 | 22 | 0.44 | No | No |
| Y4 M1 | 22 | 0.44 | No | No |
| Y4 M2 | 27 | 0.54 | No | No |
| Y4 M3 | 31 | 0.62 | Yes | No |
| Y4 M4 | 36 | 0.72 | Yes | Yes |
| Total | 180 |  |  |  |

Table 22.2 — Suturex misfire complaints against the PMS plan limits, trend window (appendix to the decision request).

| No. | Decision | Rationale | Owner | Due |
| --- | --- | --- | --- | --- |
| D1 | Trend-report criterion met; file the Article 88 trend report | 3σ breach at Y4 M4, and 2 of the last 3 months beyond 2σ | Regulatory lead | Within the applicable timeline |
| D2 | Validate as a signal; open an assessment stratified by lot (before and from SX-2310) | New severity (9 leaks with reoperation) not in the risk file | Vigilance lead | Protocol in Y4 M5 |
| D3 | Review misfire severity in the risk management file | Harm "prolonged procedure" no longer covers observed events | Risk management lead | Y4 M5 |
| D4 | Confirm individual serious-incident reportability for each of the 9 leak events | Reoperation may meet the serious-incident definition (MDR Article 2(65)) | Regulatory lead | Immediately |
| D5 | Recommend an FSCA options paper; escalate to the named executive | Possible lot-linked action with supply impact | Board chair | Next urgent session |

Table 22.3 — Decision log extract, Suturex urgent session.

Dissent recorded on D2: the engineering lead wanted to wait for returned-device results before stratifying by supplier lot. The board's answer, recorded in the minutes: validating a signal does not require a known cause, and stratification costs nothing if the hypothesis is wrong.