COMPANION TOOLKIT · TOOL T-22

# Cross-Document Reconciliation Checklist

One controlled conclusion, asked of every document that can describe the device's safety

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Lists every document that can describe a device's safety, asks the same questions of each against one controlled conclusion statement, and records the version checked, the result, the owner and the due date. |
| --- | --- |
| When to use | Before PMS deliverables (PSUR, CER, SSCP, IFU, risk management file, PMCF plan, US reports) leave the building — not after a reviewer returns them — and whenever a signal conclusion changes. |
| Chapter reference | Chapter 24, Keeping PMS Deliverables Consistent — section 24.5 (Tables 24.1–24.3). |

| Device |  | Signal record ID |  |
| --- | --- | --- | --- |
| Conclusion statement version |  | Checklist date / owner |  |

# Part 1 — Controlled conclusion statement

| Risk |  |
| --- | --- |
| Status |  |
| Failure modes |  |
| Key estimate (with interval and data cut) |  |
| Pattern |  |
| Actions |  |
| Version and approval date |  |

# Part 2 — Reconciliation map

| Document (version) | Current wording | Required change | Owner |
| --- | --- | --- | --- |
| PSUR |  |  |  |
| CER |  |  |  |
| SSCP |  |  |  |
| IFU |  |  |  |
| Risk management file |  |  |  |
| PMCF plan |  |  |  |
| US MDR coding / Part 806 assessment |  |  |  |

# Part 3 — Checklist

| Check | Version checked | Result | Owner | Status / due date |
| --- | --- | --- | --- | --- |
| Controlled conclusion statement exists and is versioned |  |  |  |  |
| PSUR status term matches the statement |  |  |  |  |
| CER includes the conclusion and the key comparison |  |  |  |  |
| PMCF plan covers the failure modes and any design change |  |  |  |  |
| SSCP status matches the statement |  |  |  |  |
| IFU warnings and follow-up match the advisory |  |  |  |  |
| Risk management file carries all modes and the revised estimate |  |  |  |  |
| Related CAPAs cite the single signal record |  |  |  |  |
| US MDR coding and Part 806 assessment match the modes and the advisory |  |  |  |  |
| One estimate used everywhere |  |  |  |  |

An "Open" line on this checklist is a finding the company made itself. The same line in a notified body report is a finding someone else made.

### Review

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Vigilance lead |  |  |  |
| Regulatory affairs lead |  |  |  |
| Clinical evaluation lead |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Kestran defibrillator lead, end of Year 6 (PSUR said "identified risk"; CER absent; SSCP "under investigation"; IFU absent).

| Element | Entry |
| --- | --- |
| Risk | Kestran lead fracture |
| Status | Confirmed: identified risk |
| Failure modes | Early fracture just distal to the header (7); wear-out fracture of the conductor at the subclavian crossing (41) |
| Key estimate | 5-year Kaplan–Meier fracture-free survival 97.7% (95% CI 97.1–98.4); predecessor lead 99.2%; data cut end of Year 6 |
| Pattern | Annual hazard rises from 0.08% in year 1 to 1.38% in year 5 (wear-out) |
| Actions | Physician advisory (no prophylactic extraction; remote monitoring with impedance alerts on; follow-up every 3 months); header strain-relief design change; single signal record linking related CAPAs |

Table 24.1 — Kestran controlled conclusion statement.

| Document | Current wording | Required change | Owner |
| --- | --- | --- | --- |
| PSUR | "Identified risk" | Keep; add both modes, the estimate, the data cut and the action status | Vigilance lead |
| CER | Absent | Add the conclusion, the survival comparison and the benefit–risk review; update the PMCF plan | Clinical evaluation lead |
| SSCP | "Under investigation" | Change to identified risk, in plain language, with the monitoring advice | Regulatory affairs lead |
| IFU | Absent | Add a fracture warning and the follow-up advice through labeling change control | Regulatory affairs and labeling |

Table 24.2 — Kestran reconciliation map: four documents against one conclusion.

| Check | Kestran result (end of Year 6) | Owner | Status |
| --- | --- | --- | --- |
| Controlled conclusion statement exists and is versioned | Yes (Table 24.1) | Vigilance lead | Done |
| PSUR status term matches the statement | Yes, "identified risk"; modes and estimate to be added | Vigilance lead | In progress |
| CER includes the conclusion and survival comparison | No | Clinical evaluation lead | Open |
| PMCF plan covers header fractures and the design change | Not yet reviewed | Clinical evaluation lead | Open |
| SSCP status matches the statement | No, "under investigation" | Regulatory affairs lead | Open |
| IFU warnings and follow-up match the advisory | No, absent | Regulatory affairs and labeling | Open |
| Risk management file carries both modes and the revised estimate | To be confirmed | Risk management lead | Open |
| Related CAPAs cite the single signal record | To be confirmed | Quality lead | Open |
| US MDR coding and Part 806 assessment match the two modes and the advisory | To be confirmed | Regulatory affairs lead | Open |
| One estimate used everywhere (Kaplan–Meier 2.3% at 5 years) | To be checked in each document | Vigilance lead | Open |

Table 24.3 — Tool T-22 extract, completed for Kestran.