COMPANION TOOLKIT · TOOL T-23

# Audit and Inspection Readiness Checklist

Self-tests that trace one complaint through everything it touched — run before an outsider does

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Turns the inspection evidence map into tests a team can run on itself. Each line names a requirement, the record that proves it, and the result of the last self-check. Includes a mock-audit exercise and a CAPA structure for failures of the signal process itself. |
| --- | --- |
| When to use | Quarterly or on the internal audit schedule; after any process change; after an inspection finding (with a look-back). Readiness is a standing condition, not a week of preparation. |
| Chapter reference | Chapter 25, Quality, Audits and Inspection Readiness — section 25.5 (Tables 25.1 and 25.2). |

| Scope (products, sites) |  | Self-check date |  |
| --- | --- | --- | --- |
| Lead |  | Previous check (date / reference) |  |

# Part 1 — Readiness tests

| Test | Source requirement | Evidence to pull | Last result | Owner |
| --- | --- | --- | --- | --- |
| Every closed complaint has a reportability decision with reason | ISO 13485 8.2.2; 21 CFR 820.35; 803.18 | Sample of closed complaints |  |  |
| "Not reportable" decisions on complaints mentioning a health effect have a second reviewer | Company SOP | Sample of closed health-effect complaints |  |  |
| Trend analysis run as the PMS plan specifies | MDR Art 84, Annex III | Trend reports vs plan schedule |  |  |
| Coding uses IMDRF terms on separate problem and health-effect axes | Company SOP | Coding QC log |  |  |
| Signal records opened when the SOP says they must be | Company SOP | Signal log vs trigger events |  |  |
| Serious-incident decisions and deadlines documented | MDR Art 87; 21 CFR Part 803 | Incident assessment records |  |  |
| Signal conclusions feed the risk management file | ISO 14971 clause 10 | Risk file update records |  |  |
| Supplier change notification in quality agreements | ISO 13485 purchasing controls | Agreement register |  |  |
| Distributor complaint-transfer timelines defined and met | Company agreements | Distribution/vigilance agreements; transfer log |  |  |
| Signal process CAPA includes a look-back | ISO 13485 8.5.2 | CAPA file |  |  |
| Decision-quality sample run monthly; disagreement rate tracked | Company SOP | Sampling records |  |  |
| Signal system health reported to management review | Company SOP | Management review minutes |  |  |
| Information requests can be answered quickly (GB: within 3 working days when MHRA requests it during an investigation) | GB PMS regulations | Retrieval drill log |  |  |

Result examples: Pass; Pass after CAPA; Partial (with scope); Fail. Test content, not presence: a filled decision field that says only "no MDR required" is a fail.

# Part 2 — Mock-audit exercise

Pick one complaint at random from last quarter. Without warning the team, ask them to produce, within two hours, every record the complaint touched. Time the exercise and note every gap. Repeat it with a complaint that mentions injury.

| Record the complaint touched | Produced? (Y/N) | Time | Gap found |
| --- | --- | --- | --- |
| Coding (IMDRF, both axes) |  |  |  |
| Reportability decision and reason (each jurisdiction) |  |  |  |
| Trend inclusion |  |  |  |
| Signal record (if any) |  |  |  |
| Risk file link |  |  |  |
| CAPA link |  |  |  |

| Complaint ID (random) |  | Complaint ID (injury) |  |
| --- | --- | --- | --- |
| Start / finish time |  | Facilitator |  |

# Part 3 — Answering in the room

☐  Back-room team retrieves records; front-room lead presents them.

☐  Answers are short and exact; when a record does not exist, say so and do not improvise.

☐  Every request and every document provided is logged.

| Request no. | Time | Request | Document provided (ID) | Provided by |
| --- | --- | --- | --- | --- |

# Part 4 — CAPA for the signal process itself

| CAPA element | Action | Owner | Evidence |
| --- | --- | --- | --- |
| Correction |  |  |  |
| System fix |  |  |  |
| Second review |  |  |  |
| Retrospective scope (look-back) |  |  |  |
| Effectiveness (criterion written before sampling) |  |  |  |

### Review

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Internal audit lead |  |  |  |
| Vigilance lead |  |  |  |
| Quality/CAPA owner |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. Neoflo central venous catheter: inspection in Year 5 found 140 complaints closed as "no MDR required" without a written rationale, 11 of which described patient injury. Mock audit after the CAPA.

| Test | Source requirement | Evidence to pull | Last result | Owner |
| --- | --- | --- | --- | --- |
| Every closed complaint has a reportability decision with reason | ISO 13485 8.2.2; 21 CFR 820.35; 803.18 | Sample of closed complaints | Pass after CAPA | Complaint-handling lead |
| Trend analysis run as the PMS plan specifies | MDR Art 84, Annex III | Trend reports vs plan schedule | Pass | Vigilance lead |
| Coding uses IMDRF terms on separate problem and health-effect axes | Company SOP | Coding QC log | Partial: some product lines lag | PMS analyst |
| Supplier change notification in quality agreements | ISO 13485 purchasing controls | Agreement register | Pass after update | Supplier quality |
| Signal process CAPA includes a look-back | ISO 13485 8.5.2 | CAPA file | Pass | Quality/CAPA owner |

Table 25.2 — Extract from Tool T-23, completed during the Neoflo mock audit.

| CAPA element | Action | Owner | Evidence |
| --- | --- | --- | --- |
| Correction | Re-decide all 140 files; file any late MDRs | Vigilance lead | Review log with rationale per file |
| System fix | Mandatory rationale, decision-maker and date fields before closure | Complaint-handling lead | Validated system change record |
| Second review | Health-effect complaints closed "not reportable" need a second reviewer | Safety physician | SOP revision; training records |
| Retrospective scope | Extend look-back to other product lines using the same closure path | Quality/CAPA owner | Scope memo and results |
| Effectiveness | Sample closed complaints on a defined schedule; target: no closure without rationale | Internal audit | Sampling plan and results |

Table 25.1 — Complaint-handling CAPA after the Neoflo inspection finding. The effectiveness criterion is company-designed; sample size and frequency are set and justified in the CAPA plan.