COMPANION TOOLKIT · TOOL T-24

# KPI Dictionary and AI Use-Case Validation Checklist

Part A: one definition per metric, with its pair · Part B: a gate before any automated function enters the signal process

Signal Management in Medical Device Vigilance · Dr. Hafez Selim, MD, PhD · Selim Medical Press · © 2026

| Purpose | Part A defines each signal-management metric so that every site calculates it the same way, and pairs every speed metric with a quality metric. Part B is the checklist a team completes before an automated or AI function enters the signal process, or returns to service. |
| --- | --- |
| When to use | Part A: before any metric goes on a leadership dashboard (a metric with no entry does not go on it); review yearly and whenever a process, system or SOP changes a start or stop event. Part B: before go-live and after any suspension. |
| Chapter reference | Chapter 26, Signal Management Metrics — section 26.5 (Tables 26.1 and 26.2); Chapter 27, AI, Automation and Real-World Data — section 27.5 (Tables 27.1 and 27.2). |

# Part A — KPI dictionary

Start from the decisions the metrics must support, not from the data that is easy to extract.

| KPI | Definition (numerator ÷ denominator) | Clock or scope | Paired with | Owner |
| --- | --- | --- | --- | --- |

### Gaming test (one per KPI)

If a team wanted to improve this number without improving the work, how would it do it? If the answer is easy, tighten the definition or strengthen the pair.

| KPI | How it could be gamed | Countermeasure (definition change or pair) |
| --- | --- | --- |

### Paired dashboard

| Speed metric | Baseline → current | Paired quality metric | Baseline → current | Paired status (both in range?) |
| --- | --- | --- | --- | --- |

### Definition change log

| Date | KPI | Change to definition (start/stop event, scope) | Reason | Trend lines marked? (Y/N) |
| --- | --- | --- | --- | --- |

Mark any trend line that crosses a definition change so it is not misread as a real shift in performance.

# Part B — AI use-case validation checklist

| Function / tool name |  | Model version |  |
| --- | --- | --- | --- |
| Vendor or internal |  | Checklist date |  |

| Item | Question | Entry for this function | Status | Owner |
| --- | --- | --- | --- | --- |
| Intended use | What decision does the output feed? |  |  |  |
| Consequence of error | Which error could harm a patient or breach a deadline? |  |  |  |
| Performance by class | Is sensitivity for the high-consequence class (e.g., injury) estimated with an interval, against a stricter acceptance criterion than the overall target? |  |  |  |
| Validation sample | Is the validation set representative, and enriched for rare high-consequence cases? |  |  |  |
| Human checkpoints | Which decisions require a named human reading the source text? |  |  |  |
| Traceability and change control | Are model version, input, output and reviewer changes recorded? |  |  |  |
| Ongoing monitoring | Which subgroup-level metrics are sampled after go-live, how often, and with what trigger? |  |  |  |
| Suspension and fallback | What suspends the function, and what manual process takes over? |  |  |  |

Overall accuracy hides the errors that matter: report performance separately for each consequence class and weight errors by consequence. Acceptance criteria are company-designed and justified in the validation plan. A function does not go live, or return to service, until every item is closed.

| Sign-off | Name | Signature | Date |
| --- | --- | --- | --- |
| Vigilance lead |  |  |  |
| Quality systems lead |  |  |  |

# Example (fictional case)

All names, numbers and events below are fictional composites from the book's case dossiers. They show the expected level of detail, not a regulatory standard. "The manufacturer" (unnamed, company-wide metrics) and the company-wide NLP complaint-coding tool.

| KPI | Definition | Clock or scope | Paired with | Owner |
| --- | --- | --- | --- | --- |
| Median complaint closure | Median days, all complaints closed in period | Start: receipt by any company channel; stop: closure approval | Returned-device evaluation rate | Complaint-handling lead |
| Returned-device evaluation rate | Returned devices with a documented engineering evaluation and conclusion ÷ returned devices received in period | Visual-only inspection does not count unless the SOP justifies it for that failure mode | Median complaint closure | Product engineering lead |
| Reopened-signal rate | Signals reopened because new evidence contradicted the conclusion ÷ signals closed in the prior 12 months | Planned periodic reviews that confirm the conclusion are excluded | Median time to signal closure | Vigilance lead |
| CAPA effectiveness | CAPAs meeting their written criterion ÷ CAPAs due for effectiveness check | Criterion must be written before the check period starts | Recurrence of the same failure mode | Quality/CAPA owner |

Table 26.2 — Extract from the Tool T-24 KPI dictionary.

| Speed metric | Baseline → target year | Paired quality metric | Baseline → target year | Paired status |
| --- | --- | --- | --- | --- |
| Median complaint closure (days) | 45 → 20 | Returned-device evaluation rate | 70% → 38% | Red: speed bought with quality |
| Median time to signal closure | Tracked from next quarter | Reopened-signal rate | 4% → 15% | Red: closures overturned |

Table 26.1 — The manufacturer's paired dashboard after redesign.

| Item | Question | NLP coding tool | Status | Owner |
| --- | --- | --- | --- | --- |
| Intended use | What decision does the output feed? | IMDRF coding; reportability suggestion (suspended) | Revised | Vigilance lead |
| Consequence of error | Which error could harm a patient or breach a deadline? | Injury coded as "no health consequence" | Defined | Safety physician |
| Performance by class | Is sensitivity for injury estimated with an interval? | Not in first audit; enriched set in progress | Open | PMS analyst |
| Human checkpoints | Which decisions require a named human reading the source text? | Reportability for any health-effect mention; validation; closure | Implemented | Complaint-handling lead |
| Traceability and change control | Are model version, input, output and reviewer changes recorded? | Version logged; override rate reported monthly | Implemented | Quality systems lead |

Table 27.2 — Extract from the Tool T-24 AI use-case validation checklist.

| System | Measure | Result | 95% CI | What the headline hid |
| --- | --- | --- | --- | --- |
| NLP coding tool | Agreement, all complaints | 178/200 = 89% | 83.9–92.6 | 4 injuries coded "no health consequence" |
| NLP coding tool | Injury coded as harmless | 4/200 = 2% | 0.8–5.0 | Each is a possible missed report |
| Lumora, 3 updated sites | Sensitivity | 129/150 = 86% | 79.5–90.7 | Zero complaints |
| Lumora, other sampled sites | Sensitivity | 279/300 = 93% | 89.5–95.4 | Premarket figure was 94% (470/500) |

Table 27.1 — Two monitoring results and what the headline number concealed (1,000 complaints coded; audit of 200 found 22 miscoded, 11%).